## 4095-FCP
|  |  | Document #: | 4095-FCP
|  |  | Rev: | P
|  |  | ECO: | 14775
Information included herein is controlled by the Export Administration Regulations (EAR) and may require an export license, license exception or other approval from the appropriate U.S. Government agency before being exported from the United States, reexported, transferred or retransferred to any non-U.S. person. Diversion contrary to U.S. law is prohibited.
Glider Functional Checkout Procedure (FCP) - Estimated Time To Complete: 90 minutes
Glider Serial Number: | calanus | Glider Firmware Rev. | Glider: Version 11.08 SW311349-NFC-F-RK CRC 0xcf9de0f4 Science: 311349.08
Tested By: | GO, TM, KB
Reason For Test: | Pre-deployment San Diego
Date: | 2026-08-18
CAUTION
1. The Pumped SBE CTD should not be run dry for more than 30 seconds at a time. *Note Pumped SBE CTD should not run until a minimum conductivity is reached. 2. The glider must not be in simulation. If glider is in simulation, follow the instructions in step 1.02 to delete the simul.sim file 3. The thruster should not be operated dry/in air for prolonged periods of time. 4. Thruster maintenance should be done regularly.
NOTES
If any test results in a failure, the FCP must be restarted after the failure is corrected. An experienced technician will often use “where” command in place of “get” when appropriate.
EQUIPMENT REQUIRED:
Dockserver/Maintenance Laptop (including Freewave transceiver) |  | Green "go" plug
Clock with UTC time |  | Hand tools to remove / replace nose cone and tail cowling.
Thermometer (Optional to confirm CTD Temp reading) |  | RF Chirp (TWR P/N E-710) - optional if glider is tested through satellite or with local receiver.
Digital multimeter
Note: this procedure should only be performed by trained and qualified glider technicians. Contact glidersupport@teledyne.com to inquire about training sessions.
Indoor Tests: |  |  |  | Pass/Fail/ Result/Note
1 | 1. Verify Connection of Glider O&M Communications Kit (Freewave properly connected to Mobile Dockserver). |  |  | Pass
| 2. Start up and sign into the Dockserver or Maintenance Laptop. |  |  | Pass
| 3. Start glider terminal |  |  | Pass
On a dockserver with SFMC installed (recommended): double-click the SFMC icon on the desktop and navigate to the respective Glider Terminal Page
Alternatively, from a Linux terminal session: type "start-glider-terminal" to start glider terminal type "inspect-dockserver" to monitor Dock Server activity (optional)
| 4. If a script.xml is running, click the stop button and make sure it stays off (red) during the procedure. |  |  | Pass
| 5. Place a "RF Chirp" beeper near the fin. During the test, verify that the RF Chirp beeps approximately every 90 seconds (or use alternative method to confirm Argos transmissions). |  |  | Pass
| It is recommended that before the glider is deployed, verify that Argos hits are accessible from your Argos provider. |  |  | Pass
| 6. Insert Green Plug or Bench Power Supply to power up the Glider. |  |  | Pass
NOTE: The Strobe Light will sometimes flash when power is applied to the Glider. Flashing or not Flashing is not indicative a of problem.
| 7. Verify the Air Pump has turned on and record time. | start time = | 1:29 | Pass
| 8. Verify that the Freewave Master indicates carrier detect (CD light turns green) within 30 seconds of applying Glider power. |  |  | Pass
EAR-Controlled Technology Subject to Restrictions Contained on the Cover Page.
1.01 | 1. Stabilize glider for testing by getting into lab_mode, to do so, perform the following steps: |  |  | Pass
| Type ctrl-C on the Glider Terminal to prevent the glider from sequencing into a mission. |  |  | Pass
NOTE: If the Glider begins to run the mission "initial.mi" before accepting a ctrl-c, you must type exit reset to restart the process. Failure to do so may result in test errors due to some sensors being disabled.
| 2. At the first GliderDOS prompt, type the following commands: |  |  | Pass
| lab_mode on |  |  | Pass
| callback 30 |  |  | Pass
| zs c:\config\autoexec.mi and longterm.sta (wait for file to transfer) |  |  | Pass
| logging on (wait for response LOG FILE OPENED) |  |  | Pass
| 3. Record the name of the log file NOTE: It could take 1-2 minutes for the log file to open. | Log file name: | calanus-2026-229-0-0 (0166.0000) calanus-2026-229-1-0 (0167.0000)
1.02 | 1. Type ver and record software version (from header in front) | Software Version: | Glider: Version 11.08 SW311349-NFC-F-RK CRC 0xcf9de0f4 Science: 311349.08
| 2. Type simul? If no response, then the glider is not in simulation. If the glider is in simulation, type del c:\config\simul.sim then type exit reset and restart the FCP |  |  | Pass
| 3. Type date and verify that the time matches the current UTC time (+/- 5 minutes). If date is significantly wrong, type sync_time while outside with a clear view of the sky. If this does not result in a stable correct date over time, the coin cell batteries may need to be replaced or the Garmin GPS may need new firmware. Contact Glidersupport@Teledyne.com for support. |  |  | Pass
| 4. Verify that the Air Pump shuts off between 3 and 10 minutes for single air pump vehicles, 1 to 5 minutes for dual air pump vehicles. | elapsed time = | 3 mins | Pass
NOTE: After the Air Pump turns off, it should not turn back on until the glider is reset - doing so is an indication of a problem, with the exception being if a device comes out service. If testing in a magnetically saturated area (concrete floors, near electronics or on deck of ship the attitude sensor may come out of service – the FCP testing can continue in this circumstance. The timing of the air bladder inflation may need to be restarted).
Altimeter Check
1.03 | 1. Remove the nose dome. |  |  | Pass
| 2. Determine the type of altimeter installed. To do this, check the installed devices section of the glider's autoexec.mi, if using 'installed altimeter' move onto 1.03.3. If using 'installed altimeter_232' skip to 1.04.1. Alternatively, the 'use' command can be utilized to determine the type of altimeter installed. |  |  | Pass
| 3. Type put c_alt_time 0 (power on the altimeter) |  |  | Pass
| 4. Verify that a regular clicking sound is heard from the Altimeter Transducer. NOTE: Close proximity or direct physical contact between the transducer and your ear will be required. |  |  | Pass
| 5. Type report ++ m_altimeter_voltage |  |  | Pass
| 6. Verify that the voltage reading updates every 10-12 seconds and changes. Typically in air, m_altimeter_voltage >2.2 volts, low voltage in air maybe an indication of a problem with the altimeter board. Contact glidersupport@teledyne.com with concerns. |  |  | Pass
| 7. Type put c_alt_time -1 to stop altimeter from pinging. |  |  | Pass
| 8. Type report clearall |  |  | Pass
1.04 | 1. Type put c_alt_time 0 (power on the altimeter) |  |  | NA
| 2. Type put u_alt_debug 1 (this puts altimeter debug into verbose mode). Verify that a regular clicking sound is heard from the Altimeter Transducer |  |  | NA
| 3. Verify altimeter debug is outputting ‘altimeter_232_power_on()’ , ‘altimeter_232_measure, received '$SDDPT……’ , and ‘altimeter_232_measure() RX valid message '$SDDPT……’ lines. |  |  | NA
| 4. Type put c_alt_time -1 (power off altimeter). Wait to see ‘altimeter_232_power_off(), power removed, uart closed’ from debug output. |  |  | NA
| 5. Verify altimeter has stopped clicking |  |  | NA
| 6. Type put u_alt_debug 0 (disable verbose altimeter debug ) |  |  | NA
Leakdetect Check
1.05 | 1. Type get m_leakdetect_voltage and verify that the reading is greater than 2.3 V and not greater than 2.5 | m_leakdetect_voltage = | 2.4893 | Pass
| 2. Type get m_leakdetect_voltage_forward and verify that the reading is greater than 2.3 V and not greater than 2.5 | m_leakdetect_voltage_forward = | 2.4935 | Pass
| 3. If G3, type get m_leakdetect_voltage_science and verify that the reading is greater than 2.3 V and not greater than 2.5 | m_leakdetect_voltage_science = | 2.4923 | Pass
| 4. For Gliders equipped with digifin, type get m_digifin_leakdetect_reading. | m_digifin_leakdetect_reading = | 1023 | Pass
| 5. Verify digifin leakdetect reading is greater than 1019. If reading is less than 1019, the digifin needs factory service. Take digifin out of service (use - digifin) and continue FCP. |  |  | Pass
EAR-Controlled Technology Subject to Restrictions Contained on the Cover Page.
Coulomb Amphr Total Check
1.06 | CAUTION: Only perform step 1.06 if a new battery pack has been installed. Otherwise, proceed to step 1.07.
| 1. Only for new batteries or recently charged rechargeable batteries: type put m_coulomb_amphr_total 0 |  |  | Pass
1.07 | 1. Type get m_coulomb_amphr_total and note the value (applies to all G2 and G3 gliders, and G1 gliders with add-on coulomb counters) Note: For gliders with new batteries, verify that energy consumption while in the lab during testing in not excessive. Battery use in the lab with vary greatly as testing requirements and energy consumption of individual gliders can vary greatly. Nominally most factory and pre-mission testing can be done using 2-10 coulomb amphr per day of testing. For all vehicles and particularly with gliders with partially used batteries, verify energy remaining will support intended glider operations. | m_coulomb_amphr_total = | 170.65 | Pass
| 2. Type get f_coulomb_battery_capacity | f_coulomb_battery_capacity = | 300 | Pass
| 3. Verify the value of f_coulomb_battery_capacity correctly reflects the battery chemistry and configuration installed in the glider. |  |  | Pass
Science Check
1.08 | 1. For gliders with a science computer - Type consci |  |  | Pass
| 2. Verify that the science prompt SciDos> is displayed. |  |  | Pass
| 3. If processor is a Persistor, type dir app. If processor is an STM-32, skip to Step 5. |  |  | Pass
| 4. Verify that the result contains the file SUPERSCI.APP for SBMB1 or sbm2sci.app for SBMB2 |  |  | Pass
| 5. Type zs c:\config\proglets.dat |  |  | Pass
| 6. Type quit |  |  | Pass
| 7. Verify that control returns to the Glider Processor (GliderLAB> is displayed) |  |  | Pass
1.09 | 1. Review proglets.dat to determine which science sensors are installed in the glider. |  |  | Pass
| 2. List each of these sensors in step 1.09a-f |  |  | Pass
| 3. Locate the glider sensor information sheet (GSI) for each sensor. Recommendations for science sensor predeployment checks should be obtained from the sensor vendor. |  |  | Pass
| 4. Type loadmission sci_on.mi |  |  | Pass
| 5. Verify science sensors are reporting as expected and that oddities/warnings/errors are not being generated by science_super |  |  | Pass
| 6. Type loadmission sci_off.mi |  |  | Pass
1.09a | Sensor name and GSI doc and rev (if applicable): CTD |  |  | Pass
If applicable, verify nominal bench output or response from sensor
1.09b | Sensor name and GSI doc and rev (if applicable): optode |  |  | Pass
If applicable, verify nominal bench output or response from sensor
1.09c | Sensor name and GSI doc and rev (if applicable): FLBBCD |  |  | Pass
If applicable, verify nominal bench output or response from sensor
1.09d | Sensor name and GSI doc and rev (if applicable): Solocam (not plugged in) |  |  | NA
If applicable, verify nominal bench output or response from sensor
1.09e | Sensor name and GSI doc and rev (if applicable): |  |  | NA
If applicable, verify nominal bench output or response from sensor
1.09f | Sensor name and GSI doc and rev (if applicable): |  |  | NA
If applicable, verify nominal bench output or response from sensor
Pressure Transducer Check
1.1 | 1. Type report ++ m_depth |  |  | Pass
| 2. Verify m_depth is ±0.5m (if m_depth is greater than ±0.5 m, type zero_ocean_pressure and try again) |  |  | Pass
| 3. Type report clearall |  |  | Pass
Strobe Light Check
1.11 | 1. Type strobe on (if Installed) |  |  | Pass
| 2. Verify that the strobe light flashes periodically. |  |  | Pass
| 3. Type strobe off |  |  | Pass
Vacuum Check
1.12 | Note – vacuum is significantly affected by temperature and glider cc pump position. Testing below will vary significantly from lab to operational conditions
| 1. Type get m_vacuum (vacuum will increase when the air bladder is inflated) |  |  | Pass
| 2. Verify that the vacuum reading is greater than 7"Hg (Hydraulic Pump vehicles) or 6"Hg (Piston Pump Vehicles) | m_vacuum = | 8.947330 inHg | Pass
1.13 | 1. Type ballast and wait for the response “The Air Pump is off” (this may take several minutes for a deep glider) |  |  | Pass
| 2. Type report clearall |  |  | Pass
| 3. Type get m_vacuum |  |  | Pass
| 4. Confirm that the value has decreased by 1.75" ± 0.5"Hg from step 1.12 and that the value is either above 5.5"Hg (Hydraulic Pump vehicles) or 4.5"Hg (Piston Pump Vehicles). | m_vacuum = | 7.250033 inHg | Pass
EAR-Controlled Technology Subject to Restrictions Contained on the Cover Page.
Anode Check
1.14 | NOTE: Crusty residue may need to be removed from anode to get a good connection. If excessively used the anode should be replaced.
| 1. Remove the tail cowling. |  |  | Pass
| 2. Aft Anode: Probe between the aft anode and the Screws attaching Ejection Weight Tube using a Digital Multimeter on ohms Ω setting. |  |  | Pass
| 3. Verify that resistance is less than 10 ohms. |  | resistance < 10 Ω? | Pass
| 4. Forward Anode: Probe between the forward Anode and Pump Flange Screws using a Digital Multimeter on ohms setting. Note, Forward anode readings may vary greatly between gliders. On G3 gliders, it is common not to get a good ground on the forward anode - this is not an indication of failure, the aft anode is sufficient. |  |  | Pass
| 5. This measurement may need to be done to the aft boom socket headed screws |  | resistance = | Pass
Wiggle Test
1.16 | 1. At the prompt type wiggle on and let the glider wiggle for 10 minutes Note that it is optionally acceptable for experienced users to start the wiggle test at 1.03 and complete other testing in parallel. If this method is used wiggle test must be complete before step 1.12 is started. |  |  | Pass
NOTE: It is not unusual for the Pitch Motor to produce warnings on startup after being idle or after transportation. If this happens, stop the wiggle, put the device back into service if necessary, and restart the wiggle.
| It is recommended that the buoyancy pump position and battery position is printed to the screen to verify that the glider is performing the full range of motion on each. (report ++ m_de_oil_vol for deep pumps; report ++ m_ballast_pumped for shallow pumps; report ++ m_battpos for battery position) |  |  | Pass
| 2. Verify that no motor driver warnings or errors are displayed on the Glider Terminal. (Operator can type use to monitor errors/warnings/oddities in the device menu) |  |  | Pass
| 3. Verify that the Digifin evenly moves from Port to Starboard. |  |  | Pass
| 4. Listen to verify that motors are not binding or stalling. (optional) |  |  | Pass
1.17 | 1. Type get m_tot_num_inflections |  |  | Pass
| 2. Verify that total is less than 20,000 inflections, and meets next mission requirements. Contact Glidersupport@teledyne.com or review manual for service recommendations | m_tot_num_inflections = | 2849 | Pass
1.18 | After wiggling for 10 minutes, type wiggle off |  |  | Pass
| Type report clearall if pump and/or battery position is being reported. |  |  | Pass
Thruster Check
1.19 | NOTE: only perform step 1.19 if thruster is installed. Note that thruster should not be operated in air for more than 30 seconds to prevent overheating.
| 1. Make sure that thruster blades are not obstructed. |  |  | NA
| 2. Type report ++ m_thruster_current |  |  | NA
| 3. Type put c_thruster_on 30 |  |  | NA
| 4. Verify that thruster spins clockwise when viewed from the aft end and that m_thruster_current updates regularly. |  |  | NA
| 5. Type put c_thruster_on 0 to turn off the thruster. |  |  | NA
| 5. Type report clearall |  |  | NA
CAUTION: Ensure Glider is dry prior to performing step 1.20 and 1.21. The presence of salt water during these steps increases the potential for accidental initiation of the recovery system release and glider weight release.
Recovery System Check
1.2 | NOTE: only perform step 1.20 if nose release recovery system is installed
| 1. Disconnect the supply lead to the recovery system at the Mecca connector. |  |  | NA
| 2. Connect the Digital Voltmeter between the Supply Lead and the forward Anode. Note: If forward anode is not providing a strong reading, probe between the Supply Lead and aft anode. |  |  | NA
| 3. Type put c_recovery_on 1 |  |  | NA
| 4. Verify that the voltage is at least 5 volts. |  | Voltage > 5V? | NA
| 5. Type put c_recovery_on 0 and wait for verification that the glider has accepted this command. |  |  | NA
| 6. Verify that the voltage in the supply lead is zero and reconnect the Recovery System Supply Lead. |  |  | NA
Ejection Weight Check
1.21 | 1. Type lab_mode off |  |  | Pass
| 2. Disconnect the Supply Lead to the Drop Weight at the Mecca Connector. |  |  | Pass
| 3. Connect the Digital Voltmeter between the supply lead and aft anode assembly or the screws attaching the ejection weight tube. |  |  | Pass
| 4. Type put c_weight_drop 1 |  |  | Pass
Respond with 'y' to confirm
| 5. Verify that the voltage is at least 5 volts. |  | Voltage > 5V? | Pass
| 6. Type put c_weight_drop 0 (Glider Software Release 7.15 and higher) - verify that the glider accepts this command |  |  | Pass
| 7. Type put m_weight_drop 0 (Glider Software Release 7.15 and higher) - verify that the glider accepts this command |  |  | Pass
| 8. Type logging off and wait for response "LOG FILE CLOSED" |  |  | Pass
| 9. If processor is a Persistor, type exit pico. If processor is an STM-32, type exit shell. |  |  | Pass
| 10. Verify that the voltage is 0, and reconnect the Drop Weight Supply Lead. |  |  | Pass
1.22 | Verify the recovery system burn wire (if present) and drop weight are plugged in. Install the Nose Cone and Tail Cover.
EAR-Controlled Technology Subject to Restrictions Contained on the Cover Page.
Outdoor Tests: |  |  |  | Pass/Fail/ Result/Note
2 | 1. Place the Glider outdoors on a level surface with an unobstructed view of the sky or be prepared to move the vehicle to this location during step 2.03. |  |  | Pass
| 2. Insert the green power plug. Note that outdoor testing MUST be done on battery power not a power supply. |  |  | Pass
2.01 | 1. Stabilize glider for testing by getting into lab_mode, to do so, perform the following steps: |  |  | Pass
| Type ctrl-C on the Glider Terminal to prevent the glider from sequencing into a mission. |  |  | Pass
| 3. At the Gliderdos prompt, Type the following commands: |  |  | Pass
| lab_mode on |  |  | Pass
| callback 30 |  |  | Pass
| logging on (wait for response LOG FILE OPENED) |  |  | Pass
| 4. Note the name of the log file NOTE: It will take 1-2 minutes for the log file to open. | Log file name: | calanus-2026-230-0-0 (0168.0000)
Attitude Sensor Check
2.02 | 1. Type report ++ m_heading m_pitch m_roll |  |  | Pass
| 2. Verify that the sensor updates periodically and m_pitch and m_roll are 0 (+/- 0.18 radian). |  |  | Pass
| 3. Rotate the glider cart and verify that m_heading changes and reads sensible heading values (Rotate glider 90° at four different points) |  |  | Pass
| 4. Type report clearall |  |  | Pass
GPS Check
2.03 | 1. Type put c_gps_on 3. The glider should begin outputting NMEA GPS strings |  |  | Pass
| 2. Verify that the GPS gets fixes within 2 minutes. |  |  | Pass
In the following example string, the highlighted A indicates a successful GPS fix. This character should turn from a V to an A when the glider begins to receive good fixes. It is not unusual for an operator to not see the V output as fixes are often acquired very quickly.
gps_diag(2)cycle#:538¦GPRMC,161908,A,5958.3032,N, 7000.5568,W,0.000,343.9,190808,0.3,W¦
| 3. Type put c_gps_on 1. The glider should stop outputting the NMEA GPS strings but the GPS will still be on. |  |  | Pass
Iridium Check
2.04 | 1. On the Glider Terminal, type callback 0 0 |  |  | Pass
| 2. Verify that the Iridium call completes successfully on primary number |  |  | Pass
| 3. Type callback 0 1 |  |  | Pass
| 4. Verify that the Iridium call completes successfully on alternate number |  |  | Pass
| 5. Type callback 30 |  |  | Pass
| Note: It is recommend that prior to deployment the operator checks the functionality of the iridium dockserver that will be used to pilot over the course of the deployment (for both primary and alternate dockserver, if they are different) - this includes verifying operator access to glider terminal page, successful connection of glider to terminal page, ability to issue commands to glider through the iridium dockserver. |  |  | Pass
EAR-Controlled Technology Subject to Restrictions Contained on the Cover Page.
Battery Check
2.05 | 1. Type get m_battery |  |  | Pass
| 2. Verify that the reading is 9.7-12.5 volts (lithium 3s), 12-16 volts (alkaline, lithium 4s and rechargeable batteries) | m_battery = | 16.545617 volts | Pass
2.06 | 1. Ensure that motors are not moving, and the air pump is off. |  |  | Pass
| 2. Type put c_argos_on 0 |  |  | Pass
| 3. Type get m_coulomb_current |  |  | Pass
| 4. Verify that the reading is between 0.086 and 0.288 amps. | m_coulomb_current = | 0.114686 amp | Pass
2.07 | 1. If G3 or vehicle has BMS (battery management system), verify all packs are contributing current. Recommended procedure: report ++ m_bms_pitch_current m_bms_aft_current m_bms_ebay_current Record "no load" values. Note: many gliders do not have ebay or energy bay packs, mark NA if the glider is not an extended energy bay glider. | no load m_bms_pitch_current = | 0.021875 amp | Pass
|  | no load m_bms_aft_current = | 0.021875 amp | Pass
|  | no load m_bms_ebay_current = | 0.033438 amp | Pass
| 2. Check current under load, this can be achieved by exercising the pump or thruster. If thruster is installed: Type put c_thruster_on 30 Verify increased current from no load conditions. Type put c_thruster_on 0 |  |  | NA
| If thruster is not installed: Retract pump using a put c_de_oil_vol or c_ballast_pumped command. Extend pump using a put c_de_oil_vol or c_ballast_pumped command. Verify increased current from no load conditions. |  |  | Pass
| 3. If G3 or vehicle has BMS (battery management system), type get m_bms_emergency_battery_voltage, verify emergency battery is reading voltage | m_bms_emergency_battery_voltage = | 12.455320 volts | Pass
2.08 | 1. Type logging off. Wait for response "LOG FILE CLOSED" |  |  | Pass
Status Mission
2.09 | 1. Type run status.mi. Wait for normal completion. Note this mission will express a number of “errors” while waiting for sensors that expect the vehicle to be in the water. These are not a concern as long as the mission completes normally. |  |  | Pass
| 2. Type callback 30 after mission completion. |  |  | Pass
House Keeping and Glider Storage
2.1 | 1. Type send *.* This command may need to be issued multiple time to purge all data on the vehicle. Repeat until the glider responds that it is sending less than 10 files on both science and flight computers. |  |  | Pass
| 2. Save the Data Files and Dock Server Log Files as a part of the Test Record |  |  | Pass
Note where files have been saved:
| 3. (Option) If desired, type dellog all to delete all logs. Ensure data is backed up before issuing command. Recommended before long missions |  |  | Pass
2.11 | 1. If glider is slated for deployment, verify glider is set to boot the application on start up by typing boot |  |  | Pass
| 2. Fully retract buoyancy pump for storage or shipment. (using a put c_de_oil_vol or c_ballast_pumped command) |  |  | Pass
| 3. (Option) type report ++ m_de_oil_vol to monitor retraction on deep pump. |  |  | Pass
| 4. Type put c_air_pump 0 to deflate the air bladder for storage |  |  | Pass
2.12 | 1. Type exit, remove power when prompted and properly store vehicle.
Signature: | Greta O'Dea, Kourtney Burger | Date: | 2026-08-19
EAR-Controlled Technology Subject to Restrictions Contained on the Cover Page.
REV | DESCRIPTION | ECO# | BY | DATE
A | Initial release | --- | --- | ---
B | Format changes, science sensor testing moved to standalone GSI docs, several tolerances tightened, wording clarified, added notice | 449 | BH, BS, CD, BA | 2012-03-27
C | Added steps for confirmation of BAM calibration values | 624 | BS | 2012-08-01
D | Removed steps for confirmation of BAM calibration values; minor edits for accuracy; added fields for noted values | 624 | BS | 2012-08-01
E | Added compass four point check. | 773 | CS | 2013-05-02
F | Added coulomb count verification; inserted put c_alt_time -1 to stop altimeter from pinging | 791 | CS | 2013-06-18
G | Rearranged steps to streamline process, added steps 1.18, 1.20.6, 1.20.7 | 12698 | BS | 2014-06-19
H | Insertion of step 1.15 | 13404 | JS | 2016-06-01
J | Added cautions 3-4, revised 1.04, 1.06, 1.07, 1.08a-f, 1.11, 1.12, 1.13, 1.19, 2.01, 2.08, 2.10 | 13994 | LG | 2018-07-12
K | Added glider Firmware Rev. to Page 1, added 3. to 1.02 | 14007 | LG | 2018-08-07
L | Clarifications, corrections and BMS sensor names | 14243 | BA | 2019-09-20
M | Step 2.11.2 added location to save functional checkout files | 14261 | GSD | 2019-10-01
N | Removed LBS steps and added clarifications on new processor | 14348 | JG | 2020-05-13
p | Cleaned up and clarified, added altimeter_232 & emergency battery check | 14775 | CG, KB | 2022-05-14
!